Back

International Journal of Cosmetic Science / Volume 32, Issue 1 / p. 35-46

Full Access

### The COLIPA in vitro UVA method: a standard and reproducible measure of sunscreen UVA protection

P. J. Matts, V. Alard, M. W. Brown, L. Ferrero, H. Gers-Barlag, N. Issachar, D. Moyal, R. Wolber

First published: 13 January 2010

https://doi.org/10.1111/j.1468-2494.2009.00542.x

Citations: 60

 Paul Matts, Procter & Gamble, London Innovation Centre, Egham, Surrey, TW20 9NW, UK. Tel.: +44 1784 474454; fax: +44 1784 474508; e-mail: matts.pj@pg.com

Synopsis

EN	FR

There is a continuing need to measure and communicate reliably the UVA protection offered by commercial sunscreens. To that end, the COLIPA (European Cosmetics Trade Association) ‘In Vitro Sun Protection Methods’ group has developed a new in vitro method for measuring UVA protection in a standardized, reproducible manner. The method is based on in vitro UV substrate spectrophotometry and convolution of resulting absorbance data with the action spectrum for the in vivo Persistent Pigment Darkening (PPD) endpoint to provide an in vitro UVA protection factor (UVAPF) which is correlated with an in vivo measure. This method has been published as a COLIPA guideline, used currently in European geographies for testing and labelling sunscreen products.

This article summarizes two ‘ring’ studies, involving eight separate testing laboratories, which both defined critical parameters for the method and validated it. In Ring Study 1, eight laboratories tested the in vitro UV transmission of a total of 24 sunscreens and, from the data, a unit dose of UVA (D0 of 1.2	J	cm−2) was defined to provide a single irradiation step which, by taking into account potential sunscreen photo-instability, gave the closest

agreement with in vivo UVAPF values. In Ring Study 2, eight laboratories tested the in vitro UV

transmission of a total of 13 sunscreens using this single irradiation step and established a

Back

very good correlation (r2	=	0.83; slope	=	0.84, P	<	0.0001) between resulting in vitro UVAPF values and corresponding values derived from the in vivo PPD method. This new method, therefore, can be used to provide a reliable in vitro metric to describe and label UVA efficacy in sunscreen products, in line with the EU Commission recommendation 2006/247/EC.

# Introduction

While sunlight provides genuine benefits to human health and well-being, there is no longer any doubt that wavelengths in the ultraviolet radiation (UVR) waveband of the solar spectrum, comprising both ultraviolet B (UVB; 290–320	nm) and ultraviolet A (UVA; 320–400	nm) radiation, represent a significant risk in the etiology of carcinogenesis and photoageing [1, 2]. While there is consensus that UVB wavelengths are responsible for the majority of the deleterious effects of UVR exposure, both acute and chronic, research in recent years has provided an accumulation of evidence for a role by UVA wavelengths also. For example, solar UVR climatology demonstrates the disproportionate ground-level dose of UVA vs. UVB [3], UVA wavelengths are scattered to a greater depth in human skin [4], UVA can promote generation of reactive oxygen species which are known to damage lipid, proteins and DNA [5], UVA is a known agent of immune suppression [6] and is a suspect in cutaneous carcinogenesis [1, 7].

In recent years, therefore, there have been concerted efforts around the globe to incorporate UVA protection into commercial sunscreens, to develop methods to measure the photoprotection afforded by these products and clear, intuitive means to communicate such protection to consumers via labelling. The COLIPA (European Cosmetics Trade Association) ‘In Vitro Sun Protection Methods’ group was given the remit to develop an in vitro measure of protection from UVA wavelengths, correlated with an in vivo measure of the same. The result is the now-published COLIPA technical guideline, ‘Method for the in vitro determination of UVA protection provided by sunscreen products’ [8].

The test is based on the measurement of UV radiation (UVR) transmission through a thin film of sunscreen sample spread on a UVR-transparent roughened substrate, before and after exposure to a controlled dose of UVR from a defined source of solar-simulated radiation (SSR). By convoluting the ensuing transmission data with the action spectrum for in vivo Persistent Pigment Darkening (PPD; an in vivo UVA endpoint used by the Japan Cosmetic Industry Association [9, 10] for testing and labelling sunscreen products) and with the emission spectrum of the filtered solar simulator used for the in vivo PPD test, an in vitro UVA-Protection Factor (UVAPF) is provided, which is correlated with its corresponding in vivo PPD value.

In developing this method, the COLIPA In Vitro Sun Protection Methods group had to overcome

two significant issues. First, there is an inherent lack of both intra-laboratory repeatability and

Back

inter-laboratory reproducibility in the amplitude of absolute in vitro UV protection factors (e.g., in vitro estimates of in vivo SPF), principally because of the difficulties in achieving homogeneous distributions of sunscreen samples across the substrate. Taking a lead from the work of Wendel et al. [11], this issue was addressed by adjusting mathematically each set of sunscreen absorbance data (both before and after UVR exposure) by a correction factor coefficient, C. This coefficient is determined iteratively from the non-exposed sample’s absorbance data (calculated from the original transmission data) such that the calculated in vitro SPF value is now matched to the sunscreen’s labelled (i.e., in vivo) SPF.

Secondly, to account for potential photo-instability of the sunscreen sample (which would be a potentially significant factor in in vivo testing), the In Vitro Sun Protection Methods group developed an irradiation step in which the sample is exposed to a dose of SSR proportional to the initial UVA protection factor UVAPF0, calculated from the adjusted absorbance data of the pre-SSR-exposed sample. For each sunscreen sample, the total irradiation dose delivered to it is calculated by multiplying UVAPF0 by a unit dose D0. The final UVA protection factor, UVAPF, is calculated from the C-adjusted absorbance data of the SSR-exposed sample.

As In Vitro Sun Protection Methods group is formed of industry members actively involved in the development and testing of sunscreens, it is able to design and execute technical ‘ring studies’, in which up to eight separate laboratories participate in running the same protocol. Ring studies are an essential part of the development of pan-geography test methodologies as they test both the ‘repeatability’ (i.e. the utility of a method in returning the same results in the same laboratory, with the same operator and equipment, etc.) and the ‘reproducibility’ (i.e. the utility of a method in returning the same results in different laboratories, with different operators and equipment, etc.) of the method. By comparing these data with known reference standards (in this case, in vivo PPD), the ‘accuracy’ of the method can be determined also.

This article, therefore, describes and summarizes the methods and results from the two ring studies critical in the development of the published COLIPA In Vitro UVA Test Method [8]. We demonstrate through these data that this new method can be used to provide a reliable in vitro metric to describe and label UVA efficacy in sunscreen products, in line with the EU Commission recommendation 2006/247/EC [12].

# Materials and methods

## Materials

### UV spectrophotometer

Test laboratories used a variety of different UV spectrophotometers to conduct the

transmission measurements, but it was ensured that they all conformed to minimum

Back

requirements, as outlined below (and as reflected in the published COLIPA In Vitro UVA Method [8]).

All UV spectrophotometers used spanned the primary waveband of interest, 290–400	nm, with a spectral resolution of at least 1	nm. Device detectors were capable of collecting both the direct and diffuse portions of UVR transmitted through the roughened PMMA substrate, either with or without applied sunscreen. The wavelength accuracy of the devices was >

±0.5	nm (as checked using a mercury spectral standard lamp or a xenon lamp with a specially doped filter) and the dynamic range of the device detectors was at least 2.2 absorbance units. The maximum measured absorbance was <90% of the dynamic range of the device used.

Lamp sources used by the spectrophotometers emitted continuous radiation with no peaks within the 290–400	nm waveband and total dose was low (<0.2	J	cm−2 per measurement cycle), so that photo-stability of the sunscreen did not become a factor during spectrophotometric measurement. Xenon flash lamps were, therefore, an ideal solution and used in many of the devices used in the studies reported herein (e.g. Labsphere 1000S).

### Monitoring of the UV spectrophotometer

UV spectrophotometers were tested at regular intervals (on a monthly basis) by the measurement of reference samples. A two-fold test was performed as recommended:

- Monitoring the instrument’s response and dynamic range with standard PMMA plates (see Appendix IIA in the published COLIPA In Vitro UVA Method [8]).

- Checking wavelength accuracy with an approved standard material (e.g. holmium perchlorate, as recommended in Appendix IIB of the published COLIPA In Vitro UVA Method [8]).

### Source for SSR irradiation

The spectral irradiance of the artificial UVR source (at the sample plane) that was used for irradiation of product samples was as similar as possible to the irradiance at ground level under a standard zenith sun as defined by COLIPA (2006) [13] or in DIN67501 (1999) [14]. Target UV irradiance was set within the following acceptance limits (measured at the sample plane):

Total UV irradiance (290–400	nm) 50–140	W	m−2

Irradiance ratio of UVA (320–400	nm) to UVB (290–320	nm) 8–22

Devices used had the ability to maintain a temperature of <40°C at sample level. An example of a device meeting these specifications (and indeed, all laboratories taking part in these studies used these devices) is the Atlas Suntest™ insolator (Atlas Material Testing Technology GmbH, Linsengericht, Germany), type CPS/CPS+ or XLS/XLS+, fitted with the UV short cut-off filter and IR dichroic mirror (Atlas part references: 56052371 and 56052059 respectively).

### Monitoring of the SSR source

The emission of the SSR source was checked by a suitably qualified expert using a calibrated spectroradiometer for compliance with the given acceptance limits. On an on-going basis, the SSR source emission was also monitored using a radiometer. All spectroradiometers and radiometers were calibrated according to COLIPA recommendations in the Guideline ‘Monitoring of UV light sources’ (2007). The importance of this monitoring cannot be over-emphasized as it provides a coefficient of correction between radiometry and spectroradiometry that helps ensure that all laboratories are applying the same UVR dose [15].

### Substrate

The test sunscreen was applied to a substrate which was UVR-transparent, non-fluorescent, photo-stable and inert to all potential sunscreen formulation ingredients. Furthermore, the substrate needed to have a roughened topography on its upper surface to reflect the textured nature of human skin. Taking these requirements into account, polymethylmethacrylate (PMMA or Plexiglas™) was chosen as a suitable test substrate. Plates made of PMMA were >16	cm² and square in shape (as per the now-published COLIPA In Vitro UVA Method [8]).

Plate roughness is a critical parameter [16] and parameters for this are described in Appendix III of the published COLIPA In Vitro UVA Method [8]. For reference, the average roughness value (Sa) of the PMMA plates (supplied by Schönberg GmbH, Hamburg) used in all studies described in this article was approximately 2	μm (Sa defined by EUR 15178 EN [17]) and they were of dimensions 50	×	50	×	2.5	mm.

To help keep absorbance of tested sunscreens below 2.0	units within the 290–400	nm waveband, an application rate of 0.75	mg	cm−2 was chosen and used in all studies reported in this article.

## Common study procedures

### Control plates

Reference plates that controlled for the substrate characteristics were produced by spreading a

few microlitres of glycerine or another appropriate UVR-transparent substance on the

Back

roughened side of the plate, using just enough to thinly coat the entire plate surface (approximately 15	μL for a 50	×	50	mm plate).

### Sample application

Sunscreen product was applied at an application rate of 0.75	mg	cm−2 to the roughened side of PMMA plates as a large number of small droplets of approximate equal volume, distributed evenly over the whole surface of the plate. Positive-displacement automatic pipettes were found to be ideal for this purpose. To check for the correct application rate, pipettes and/or plates were weighed before and after dispensing the product.

After application and check-weighing, the sunscreen product was spread immediately over the whole surface using light strokes with a naked fingertip (i.e., no finger-cot) ‘pre-saturated’ with a small amount of the product. Spreading was completed in two phases: (a) the product was first distributed over the entire plate using light pressure, in <30	s and (b) the distributed sample was then rubbed into the rough surface using stronger pressure over a period of 20–30	s.

Treated samples were then allowed to equilibrate in the dark, at ambient temperature, for at least 15	min to help facilitate formation of a stabilized product film.

### Transmission measurements through product-treated plates

The UV transmission (monochromatic absorbance data over 290–400	nm with 1	nm steps) of at least three PMMA plates, treated with product in the manner described above, was measured using a calibrated UV spectrophotometer. Each plate was measured at a number of different sites to ensure that an area of at least 2.0	cm2 was measured in total. Single spot sizes were >0.5	cm2 (so that, for example, if a spot size of 0.6	cm2 was used, then at least four separate measurements on different areas of the plate were required so that the total measurement area exceeded 2.0	cm2). As far as possible, it was ensured that exactly the same location on the plates was measured before and after SSR exposure. This was to help reduce variability in measurement because of inhomogeneity of sunscreen distribution across the plate. Practically, this re-alignment can be achieved through the use of custom templates, etc.

### Exposure to SSR

Care was taken to ensure that samples were not exposed to temperatures of >40°C during irradiation. In the Atlas Suntest™ insolators, for example, air-conditioning units or water-cooled trays were employed for this purpose. PMMA plates were also fixed in place using suitable means (e.g., double-sided tape or the use of a template with wells in the surface to

## Ring study 1

The purpose of this study was to select a suitable unit dose of UVA (D0) for the SSR step (the full

dose D calculated by D0	×	UVAPF0), to take into account the important aspect of sunscreen

photo-stability and, thus, achieve the closest possible agreement between in vitro UVAPF and in vivo UVAPF values (as determined by the PPD method).

In 2004, therefore, a ring test was performed by eight laboratories (L1–L8) on 24 commercially available sunscreen products from major European geographies, comprising products containing organic filters only and a combination of both organic and mineral filters. SPF values were taken from product labelling, but in vivo UVAPF values were determined by the PPD method (using n	=	5 subjects each) across four separate contract clinical test institutes, some products being tested more than once (Table I).

Table I. In vivo UVAPF values (derived from the PPD method, n	=	5 each) reported from four separate measurement laboratories, for sunscreen products used in Ring Study 1

|  | Sunscreen sample | SPF | PPD in vivo Institute A, n = 5 | PPD in vivo Institute B, n = 5 | PPD in vivo Institute C, n = 5 | PPD in vivo Institute D, n = 5 | PPD in vivo mean of means |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | S1 | 25 | 4.1 |  |  |  | 4.1 |  |  |
|  | S2 | 40 | 4.6 |  |  |  | 4.6 |  |  |
|  | S3 | 20 | 3.8 |  |  |  | 3.8 |  |  |
|  | S4 | 20 | 11 |  |  |  | 11.0 |  |  |
|  | S5 | 20 | 15.8 |  |  |  | 15.8 |  |  |
|  | S6 | 30 | 12 |  |  |  | 12.0 |  |  |
|  | S7 | 20 | 4.4 |  |  |  | 4.4 |  |  |
|  | S8 | 20 | 3.4 |  |  |  | 3.4 |  |  |
|  | S9 | 20 | 3.4 | 2.2 |  |  | 2.8 |  |  |
|  | S10 | 20 | 4.5 | 3.2 |  |  | 3.8 |  |  |
|  | S11 | 25 | 4.0 | 5.8 |  |  | 4.9 |  |  |

|  |  |  | n | = | 5 | n | = | 5 | n | = | 5 | n | = | 5 | means |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Sunscreen	SPF	PPD in vivo	PPD in vivo	PPD in vivo	PPD in vivo	PPD in vivo	

sample	Institute A,	Institute B,	Institute C,	Institute D,	mean of

Back

The 24 sunscreens were distributed among the eight different test laboratories according to a balanced randomization scheme, such that all sunscreen products were tested by four separate laboratories. First of all, in vitro transmission measurements were made for each sunscreen product, prior to exposure to SSR, to acquire initial A0(λ) data. These data were then adjusted using a coefficient C to achieve an in vitro SPF equal to the labelled in vivo SPF (see Equations 1 and 2 in Appendix 1). This coefficient is applied to overcome the inherent lack of intra- and inter-laboratory reproducibility in the amplitude of absolute in vitro UV protection factors. Initial UVAPF0 values were then calculated (see Equation 3 in Appendix 1).

The sunscreen samples were then exposed to progressive increments of SSR using Atlas Suntest™ insolators in all laboratories. All such instruments were measured spectroradiometrically by an independent expert, (UV-Technik-Hanau). This enabled the calculation of specific exposure times at a defined irradiance for each laboratory to achieve standardized, accurate doses of UVA (see Table II for individual laboratory data). For each sunscreen sample, a nominal maximal ‘100%’ UVA dose was derived from the following

calculation, SPFlabelled	×	2.0	J	cm−2 UVA. Ideally, as many dose increments as possible

between 0% (initial) and 100% would have been applied to each sunscreen, with measurements of UVR transmission performed immediately afterwards. This, of course, was simply not practical and so each laboratory performed at least three dose increments, namely 12.5%, 25% and 50% (although some laboratories also included 75% and 100% increments). The irradiation time needed for each individual Suntest™ instrument was calculated from values found in

Table II.

Table II. Spectroradiometric data for the Atlas SuntestTM insolators used in each of the eight participating test laboratories (data for test laboratory four not available)

|  | Test site | UV (W m−2) | UV-A (W m−2) | UV-A1 (W m−2) | UV-A2 (W m−2) | UV-B (W m−2) | UV-c (mW m−2) | Time to reach 100 kJ m−2in UV-A (min) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Lab1 | 100.7 | 96.4 | 84.1 | 12.3 | 4.3 | 0.0000 | 17.28 |  |
|  | Lab 2 | 104.6 | 98.2 | 84.3 | 13.8 | 6.5 | 0.0000 | 16.98 |  |
|  | Lab 3 | 105.8 | 99.5 | 86.0 | 13.4 | 6.3 | 0.0000 | 16.75 |  |

Test	UV (W	UV-A	UV-A1	UV-A2	UV-B	UV-c	Time to reach 100

site	m−2)	(W	m−2)	(W	m−2)	(W	m−2)	(W	(mW	m−2)	kJ	m−2in UV-A (min)

Back

|  |  |  |  |  |  | m−2) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Lab 5 | 99.9 | 95.1 | 82.4 | 12.7 | 4.8 | 0.0000 | 17.52 |  |
|  | Lab 6 | 114.9 | 110.1 | 96.9 | 13.2 | 4.8 | 0.0000 | 15.14 |  |
|  | Lab 7 | 102.6 | 95.5 | 82.2 | 13.3 | 7.1 | 0.0000 | 17.45 |  |
|  | Lab 8 | 100.1 | 93.4 | 80.1 | 13.3 | 6.7 | 0.0000 | 17.85 |  |

After each dose increment, in vitro UVR transmission measurements were again performed and UVAPFi values (i.e., UVAPF values after a respective irradiation) were calculated, each adjusted by coefficient C (see Equation 4 in Appendix 1).

As only a limited number of UVAPFi values could be derived from each assay, a mathematical model describing the relationship between UVAPFi and UVR dose was therefore developed to interpolate and provide predicted values lying between measured values. A mathematical inverse exponential model was found to be very suitable to represent the variation in UVAPFi throughout continuous UVR exposure, taking into account UVAPF0:

Where, D is UV dose in J	cm−2 such that UVAPFi	=	UVAPF0 when D	=	0 and, as a

prerequisite, where UVAPFi	≥	1 such that UVAPF0– a	≥	1 when D	=	∞

Determination of the function parameters a and b was achieved for each in vitro measurement using Excel Solver (Microsoft Corporation, Redmond, Washington, USA) with the final fit to actual UVAPFi data being determined by least-square error assessment (minimum square biases). The model satisfactorily predicted actual study data as can be seen in Fig. 1 for a particular sunscreen sample where the model curve and the actual observed UVAPFi data, plotted against UV dose, are in very good agreement.

Back

#### Figure 1

Open in figure viewer

PowerPoint

Comparison of actual data (solid points) and modelled data (continuous line) from the mathematical inverse exponential model used to predict variation in UVAPFi throughout continuous UVR exposure, taking into account UVAPF0 (data selected from one representative test product).

# Ring study 2

The purpose of this study was to validate the choice of a single unit dose value of 1.2	J	cm−2 UVA and, thus, finalize the UVA In Vitro Method.

In 2004, therefore, a second ring test was performed by eight laboratories (L12–L82– note different coding and assignment to Ring Study 1) on 12 commercially available sunscreen products from major European geographies and one laboratory sample from Japan. This selection included products containing organic filters only and a combination of both organic and mineral filters. Once again, SPF values were taken from product labelling, and in vivo UVAPF values were determined by the PPD method (using n	=	10 subjects each) in a single contract clinical test institute (Table III).

Table III. In vivo UVAPF values (derived from the PPD method, n	=	10 each) reported from a single measurement laboratory, for sunscreen products used in Ring Study 2

Sunscreen sample	SPF	PPD, n	=	10	SD

| Back |  |  |  |  |
| --- | --- | --- | --- | --- |
| RRT01 | 60 | 14.2 | 1.5 |  |
| RRT02 | 60 | 12.5 | 2.6 |  |
| RRT03 | 40 | 22.3 | 3.3 |  |
| RRT04 | 60 | 13.0 | 2.2 |  |
| RRT05 | 25 | 8.8 | 0.6 |  |
| RRT06 | 30 | 7.9 | 0.8 |  |
| RRT07 | 60 | 25.5 | 3.5 |  |
| RRT08 | 25 | 10.5 | 2.0 |  |
| RRT09 | 25 | 8.0 | 1.4 |  |
| RRT10 | 25 | 10.4 | 2.9 |  |
| RRT11 | 30 | 7.6 | 1.2 |  |
| RRT12 | 15 | 7.8 | 1.2 |  |
| RRT19 | 30 | 6.6 | 1.1 |  |

The 13 sunscreens were distributed among the eight different test laboratories according to a balanced randomization scheme, such that all sunscreen products were tested by four separate laboratories. Each laboratory then tested each of the four sunscreen samples assigned to it according to the following procedure (with detailed procedure and manipulation as described above):

- Standardized sunscreen film application on roughened PMMA plates (0.75	mg	cm−2).

- Transmission measurement with a UV spectrophotometer.

- Iterative adjustment of the absorbance values by coefficient C such that in vitro SPF now equals in vivo SPF.

- Calculation of UVAPF0 from the corrected spectra.

- Determination of the irradiation dose D; D	=	UVAPF0	×	1.2	J	cm−2.

6 Irradiation of the sunscreen samples with dose D.

Back

- Calculation of UVAPF (in vitro UVAPF) from the absorbance values after irradiation, adjusted by coefficient C.

- As a final step, calculation of the ratio of labelled in vivo SPF to in vitro UVAPF.

# Results

## Ring study 1

Comparison of initial UVAPF0 values from non-irradiated sunscreen samples to corresponding known in vivo UVAPF values determined by the PPD method gave only a relatively poor

correlation of r2	=	0.3 (bilateral Pearson test) and a slope of 1.29 (Fig. 2).

#### Figure 2

Open in figure viewer

PowerPoint

Comparison of in vivo UVAPF (as derived from the PPD method) and in vitro UVAPF0 values (no irradiation) in Ring Study 1.

The inverse exponential model was able to predict an infinite number of UVAPFi values across the 0–100% dose range for all 24 sunscreen samples. The square value of the mean of the biases (i.e., the difference between measured in vivo PPD and in vitro UVAPF) of 28 predicted values were used to calculate the value of D0 where mean bias approached a value of zero (i.e., the closest prediction of actual in vivo UVAPF). The result can be seen in Fig. 3; note the

increased number of predicted value pairs as mean bias approaches zero, to help ensure an

accurate estimate of the desired unit dose D0. From these data, therefore, it was found that the

Back

best prediction of in vivo UVAPF occurred for total irradiance doses that were calculated by multiplying UVAPF0 by a unit dose D0 value of 1.2	J	cm−2.

#### Figure 3

Open in figure viewer

PowerPoint

A plot of Unit UVA Dose (J	cm−2) against the square value of the mean of the biases (i.e., the difference between measured in vivo PPD and in vitro UVAPF) of 28 predicted values (paired data derived from the inverse exponential model).

Taking this D0 value of 1.2	J	cm−2, UVAPFi values for all 24 sunscreen samples were calculated. When these were compared with their corresponding in vivo UVAPF values

determined by the PPD method (Fig. 4), a very good correlation was found (r2	=	0.75;

bilateral Pearson test) with a significant (P	<	0.0001) slope of 0.98 (indicating a high degree of accuracy in prediction of absolute in vivo UVAPF values).

Back

#### Figure 4

Open in figure viewer

PowerPoint

Comparison of in vivo UVAPF (as derived from the PPD method) and in vitro UVAPFi values

(after irradiation by a UVA dose, calculated using D0	=	1.2	J	cm−2) in Ring Study 1.

The full results of Ring Test 1 can be seen in Table IV.

Table IV. A comparison of SPF, in vivo UVAPF (derived from the PPD method) and in vitro UVAPF

values, both before and after irradiation (with a UVA dose calculated using D0	=	1.2	J cm−2), in Ring Study 1

|  | Sunscreen sample | SPF | PPD in vivo | UVAPF0in vitro | UVAPFiin vitro |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | S1 | 25 | 4.1 | 7.3 | 5.4 |  |  |
|  | S2 | 40 | 4.6 | 8.6 | 5.4 |  |  |
|  | S3 | 20 | 3.8 | 5.3 | 4.6 |  |  |
|  | S4 | 20 | 11.0 | 9.2 | 8.5 |  |  |
|  | S5 | 20 | 15.8 | 14.7 | 13.5 |  |  |
|  | S6 | 30 | 12.0 | 15.0 | 13.3 |  |  |
|  | S7 | 20 | 4.4 | 6.4 | 2.8 |  |  |

| Back |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| S8 | 20 | 3.4 | 5.2 | 4.4 |  |  |
| S9 | 20 | 2.8 | 3.2 | 3.1 |  |  |
| S10 | 20 | 3.8 | 5.2 | 5.0 |  |  |
| S11 | 25 | 4.9 | 4.1 | 4.0 |  |  |
| S12 | 20 | 8.7 | 9.5 | 8.6 |  |  |
| S13 | 12 | 2.3 | 2.8 | 2.0 |  |  |

## Ring study 2

Comparison of in vitro UVAPF values with their corresponding in vivo UVAPF values measured by

the PPD method yielded a very good correlation (Fig. 5), with r2	=	0.83 (bilateral Pearson test)

and a significant (P	<	0.0001) slope of 0.84 (indicating a good prediction of absolute in vivo UVAPF values). Furthermore, the range of in vitro UVAPF data (spanning 3.7–18.1) indicated that the method was able to measure and distinguish across a broad range of efficacy. Inter-laboratory error was low [the mean coefficient of variance (CoV) between laboratories, taking into account all samples, it was <15%].

Sunscreen sample	SPF	PPD in vivo	UVAPF0in vitro	UVAPFiin vitro	

Figure 5

Open in figure viewer

PowerPoint

Figure 6 shows UVAPF0 and UVAPF values plotted one against the other for all sunscreen samples, providing an indication of relative photo-stability. Five of the 13 sunscreen samples (RRT02, RRT06, RRT08, RRT11 and RRT19) showed minimal reduction in UVA protection (1–8% only). Inspection of the spectra of two sunscreens (RRT06 and RRT11, both with a labelled SPF of 30) indicated that they contained only a small effective amount of UVA filter. While a moderate decrease in UVA protection (17–32%) occurred for RRT01, RRT03, RRT04, RRT05, RRT07 and RRT12 when irradiated, a dramatic reduction in UVAPF (32–69%) was noted for two sunscreens in particular (RRT09 and RRT10).

#### Figure 6

Open in figure viewer

PowerPoint

Comparison of in vitro UVAPF values before and after irradiation in Ring Study 2 (mean	± SD).

Finally, Figure 7 shows the values for the ratio of labelled in vivo SPF to in vitro UVAPF for all sunscreens tested. The higher the value, the less UVA protection is afforded, relative to protection from erythemally effective wavelengths (weighted towards shorter, UVB wavelengths) [18]. Ratio values of between 2.4 and 8.2 were recorded, indicating the method

#### Figure 7

Open in figure viewer

PowerPoint

Ratio of labelled (in vivo) SPF to in vitro UVAPF values in Ring Study 2 (mean	±	SD).

Table V. A comparison of SPF, in vivo UVAPF (derived from the PPD method) and in vitro UVAPF values, both before and after irradiation, in Ring Study 2

|  | Sunscreen sample | SPF | PPD in vivo | UVAPF0in vitro | UVAPF in vitro |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | RRT1 | 60 | 14.2 | 22.5 | 16.5 |  |
|  | RRT2 | 60 | 12.5 | 13.1 | 12.6 |  |
|  | RRT3 | 40 | 22.3 | 18.9 | 16.9 |  |
|  | RRT4 | 60 | 13.0 | 18.1 | 12.6 |  |
|  | RRT5 | 25 | 8.8 | 7.4 | 4.6 |  |
|  | RRT6 | 30 | 7.9 | 4.3 | 4.3 |  |
|  | RRT7 | 60 | 25.5 | 28.0 | 19.6 |  |

Sunscreen sample	SPF	PPD in vivo	UVAPF0in vitro	UVAPF in vitro

Back

| RRT8 | 25 | 10.5 | 10.8 | 9.2 |  |
| --- | --- | --- | --- | --- | --- |
| RRT9 | 25 | 8.0 | 22.1 | 8.6 |  |
| RRT10 | 25 | 10.4 | 26.2 | 7.6 |  |
| RRT11 | 30 | 7.6 | 3.9 | 3.7 |  |
| RRT12 | 15 | 7.8 | 7.0 | 5.5 |  |
| RRT19 | 30 | 6.6 | 5.3 | 5.1 |  |

# Discussion

The two ring studies were designed both to finalize a unit dose of UVR for an irradiation step (to account for sunscreen photo-instability) and also to validate the method incorporating this step. In summary, the goals of each study were realized.

In the first ring test, it was recognized that, for some photo-instable sunscreen products, there is a progressive deterioration in the absorption of UVA wavelengths and a concomitant decrease in associated UVAPF values during irradiation with filtered UVA wavelengths in in vivo PPD testing. The measured in vivo UVAPF value is actually, thus, an integration of all the UVAPF values across the irradiation period. It was decided, therefore, to include an irradiation step into the in vitro COLIPA method to reflect this phenomenon. The challenge lay in deciding which UVA dose to use. This was approached by, first of all, taking a robust approach to calibration of the irradiation sources in all laboratories, using a single spectroradiometer and expert practitioner. These calibration data yielded, for each laboratory and device, irradiation times at a defined flux to achieve a certain unit dose of UVA radiation. In designing the ring study, an innovative approach was taken to the practical problem of modelling the progressive degradation of UVA protection (for photo-instable sunscreen products) experienced during in vivo PPD testing. A finite number of progressive total UVA doses were applied to sunscreen samples on PMMA plates (using the calibrated irradiation devices) and an inverse exponential model was used to predict UVAPFi values lying between these doses (each UVAPFi value being the product of a given unit dose of UVA, D0, and an initial pre-irradiation UVAPF0). It was found that a unit dose of 1.2	J	cm−2 of UVA, when multiplied by a sunscreen’s pre-irradiation UVAPF0, gave the most accurate prediction of in vivo UVAPF (measured by the PPD method).

The second ring study employed this defined UVA unit dose and, across eight laboratories testing 13 sunscreens of varying UVA protection and photostability, we achieved very good

repeatability (intra-laboratory CoV values for UVAPF	< 10%), reproducibility (inter-laboratory

CoV values for UVAPF	<	15%) and accuracy (between UVAPF and in vivo UVAPF values, r2

Back

=	0.82, slope	=	0.84, P	<	0.0001). Furthermore, when the results from both ring studies were combined, providing a total of 37 in vivo/in vitro comparisons and spanning a wide range

of values, a strong correlation of r2	=	0.81 with a slope of 0.78 was achieved (Fig. 8).

#### Figure 8

Open in figure viewer

PowerPoint

Comparison of in vivo UVAPF (as derived by the PPD method) and in vitro UVAPF values (after irradiation) when Ring Studies 1 and 2 are combined.

The validation of this in vitro method, which correlates with an in vivo UVA endpoint, fulfils the remit given to the In Vitro Sun Protection Methods group and provides the European industry with a repeatable, reproducible and accurate assay for measuring UVA protection.

# Acknowledgements

The authors wish to acknowledge the contributions of Sarah Meredith who provided invaluable support to the COLIPA in vitro Sun Protection Methods Group during the work described in this article.

# Appendix

Appendix 1

Equations used in the method in this article:

Calculation of the in vitro SPFin vitro

Back

(1)

where:

E(λ)	=	Erythema action spectrum (CIE 1987 [18])

I(λ)	=	Spectral irradiance received from the UV source

A 0(λ)	=	Mean monochromatic absorbance of the test product layer before UV exposure dλ	=		Wavelength step (1	nm)

## Calculation of the adjusted SPFin vitro,adj and determination of the coefficient of adjustment C

C is the coefficient of adjustment, iteratively determined to adjust the calculated in vitro SPF value to the labelled (in vivo) SPF value. It is recommended that C ranges between 0.8 and 1.2.

(2)

where:

E(λ), I(λ), A0(λ) and dλ are defined in equation (1).

## Calculation of UVAPF0

UVAPF0 is calculated for each plate individually.

2	Yaar, M. and Gilchrest, B.A. Photoageing: mechanism, prevention and therapy. Br. J. Dermatol. 157, 874–887 (2007).

CAS	PubMed	Web of Science®	Google Scholar

Back

(3)

where:

P(λ)	=	PPD action spectrum

I(λ), A0(λ), C and dλ are defined in equations (1) and (2)

## Calculation of UVAPF of plates after UV irradiation of the sample

(4)

where:

P(λ), I(λ), C and dλ are defined in equation (3).

A(λ) is the mean monochromatic absorbance of the test product layer after UV exposure.

1	De Gruijl, F.R. Photocarcinogenesis: UVA vs. UVB radiation. Skin. Pharmacol. Appl. Skin Physiol., 15, 316–320 (2002).

CAS	PubMed	Web of Science®	Google Scholar

References

3	Matts, P.J. Solar ultraviolet radiation: definitions and terminology. Dermatol. Clin. 24, 1–8 (2006).

CAS	PubMed	Web of Science®	Google Scholar

Back

- Bruls, W.A., Slaper, H., Van Der Leun, J.C. and Berrens, L. Transmission of human epidermis and stratum corneum as a function of thickness in the ultraviolet and visible wavelengths. Photochem. Photobiol. 40, 485–494 (1984).

CAS	PubMed	Web of Science®	Google Scholar

- Klotz, L.O., Holbrook, N.J. and Sies, H. UVA and singlet oxygen as inducers of cutaneous signaling events. Curr. Probl. Dermatol. 29, 95–113 (2001).

CAS	PubMed	Google Scholar

- Baron, E.D., Fourtanier, A., Compan, D., Medaisko, C., Cooper, K.D. and Stevens, S.R. High ultraviolet A protection affords greater immune protection confirming that ultraviolet A contributes to photoimmunosuppression in humans. J. Invest. Dermatol. 121, 869–875 (2003).

CAS	PubMed	Web of Science®	Google Scholar

- Bennett, D.C. Ultraviolet wavebands and melanoma initiation. Pigment Cell Melanoma Res. 21, 520–524 (2008).

CAS	PubMed	Web of Science®	Google Scholar

- COLIPA (European Cosmetics Association). Method for the in vitro determination of UVA protection provided by sunscreen products, COLIPA Guideline (2009).

Google Scholar

- Japan Cosmetic Industry Association. Measurement standards for UVA protection efficacy, Japan Cosmetic Industry Association Guideline, 1996.

Google Scholar

- Moyal, D., Chardon, A. and Kollias, N. UVA protection efficacy of sunscreens can be determined by the persistent pigment darkening (PPD) method. (Part 2). Photodermatol. Photoimmunol. Photomed. 16, 250–255 (2000).

CAS	PubMed	Web of Science®	Google Scholar

11	Wendel, V., Klette, E., Wittern, K.P. and Gers-Barlag, H. The influence of pre-irradiation on the

predictability of in vivo UVA protection with a new in vitro method. Photodermatol. Photoimmunol.

Back

Photomed. 19, 93–97 (2003).

CAS	PubMed	Web of Science®	Google Scholar

- European Commission, Enterprise and Industry, Cosmetics Commission recommendation of 22 September 2006 on the efficacy of sunscreen products and the claims made relating thereto, 2006/647/EC. Official Journal of the European Union L265, 39–43 (2007).

Google Scholar

- COLIPA (European Cosmetics Association). International Sun Protection Factor (SPF) Test Method, COLIPA Guideline (2006).

Google Scholar

- Deutsches Institut für Normung (DIN). Experimental evaluation of the protection from erythema of external sunscreen products for the human skin. Revision of DIN 67501, 9/1999 (1999).

Google Scholar

- COLIPA (European Cosmetics Association). Guidelines for monitoring UV radiation sources, COLIPA Guideline (2007).

Google Scholar

- Ferrero, L., Pissavini, M., Dehais, A., Marguerie, S. and Zastrow, L. Importance of Substrate Roughness for In Vitro Sun Protection Assessment. IFSCC Mag. 9, 97–108 (2006).

Google Scholar

- Stout, K.J., Sullivan, P.J., Dong, W.P., Mainsah, E., Lou, N., Mathia, T. and Zahouani, H. The development of methods for the characterisation of roughness in three dimensions, Report EUR 15178 EN. EC Brussels (1993).

Google Scholar

- McKinlay, A.F. and Diffey, B.L. A reference action spectrum for ultra-violet induced erythema in human skin. In: Human Exposure to Ultraviolet Radiation: Risks and Regulations ( W.F. Passchier and B.F.M.

Back

Bosnjakovich, eds), pp. 83–87. International Congress Series, Elsevier, Amsterdam (1987).

Google Scholar

Citing Literature

Download PDF

ABOUT WILEY ONLINE LIBRARY

Privacy Policy Terms of Use About Cookies

Manage Cookies Accessibility

Wiley Research DE&I Statement and Publishing Policies Developing World Access

HELP & SUPPORT

Contact Us Training and Support

DMCA & Reporting Piracy

OPPORTUNITIES

Subscription Agents Advertisers & Corporate Partners

CONNECT WITH WILEY

The Wiley Network Wiley Press Room

Copyright © 1999-2025 John Wiley & Sons, Inc or related companies. All rights reserved, including rights for text and data mining and

training of artificial intelligence technologies or similar technologies.

Back